## Slide 1: Metrological validation and operational assessment of 4-beam Wind Iris turbine-mounted Lidar
PCWG – 18th meeting
11 August 2016
Presentation by Evan Osler – evan@rnrgsystems.com
Adapted from Paul Mazoyer, Leosphere-Avent

## Slide 2: About Leosphere
Founded in 2004; independently owned
Designs, develops, manufactures, sells and services turnkey remote-sensing instruments for wind measurement and aerosol characterization
More than 750 Lidars sold in 50+ countries in several market: Wind power (70%), Meteorology, Aviation weather
Global headquarters: Orsay, France

## Slide 3: About Renewable NRG Systems
Founded in 1982.
Manufactures and distributes decision support tools for the renewable energy and broader meteorology industries.
Sells, services and supports Leosphere Lidars in North America.
Global headquarters in Hinesburg, Vermont.

## Slide 4: Agenda
4-beam Wind Iris description
Wind speed reconstruction algorithm description
Metrological validation of the 4-beam Wind Iris
Power performance testing from 4-beam Wind Iris

## Slide 5: 4-beam Wind Iris description

## Slide 6: 4-beam Wind Iris
The 4-beam Wind Iris (“Wind Iris 2”) is the new Leosphere Lidar dedicated to power curve measurement and performance optimization
  2 beams measure upward and 2 beams measure downward.
  The horizontal opening angle is 30°, similar to 2-beam Wind iris
  The vertical opening angle is 10°, for shear measurements
  New design optimized for operation, service and cost.
Two three month campaigns at operational wind farms :
  USA : Sun Edison and DNV GL. Presented at PCWG 3/25/2015. Paul Lawson, “3D Nacelle Mounted Lidar in Complex Terrain”
  Europe, France: EDF EN. Simple terrain. To be presented today!

## Slide 7: 4-beam Wind Iris calibration
In the framework of UniTTe (Uniform Turbine Testing), a procedure for 4-beam Wind Iris calibration has been designed : “White box” type.
Comparison of radial wind speed measurement with cup anemometer undertaken at RISOE test site showed high agreement.
Leosphere has also designed along with DTU an in-house calibration based on reference Lidars which is available for Windcube and may be extended to Wind Iris (“Lidar to Lidar calibration”).
GENERIC CALIBRATION PROCEDURES FOR NACELLE-BASED PROFILING LIDARS; A. Borraccino, M. Courtney, R. Wagner; DEWEK 2015

## Slide 8: Wind speed reconstruction algorithm description

## Slide 9: Wind profile is reconstructed
For each of 10 measurement distances :
Two wind speeds are measured at two different heights
The entire wind profile is then extrapolated

## Slide 10: Wind characteristics are measured
From two heights of measurement, different wind characteristics are calculated:
  Wind speed assuming log profile
  Wind direction assuming linear profile
  TI assuming log profile
Range gate | Range gate height difference [m]
#1 | 8
#2 | 13
#3 | 20
#4 | 27
#5 | 34
#6 | 40
#7 | 47
#8 | 54
#9 | 61
#10 | 67
β=5°

## Slide 11: Metrological validation of the 4-beam Wind Iris

## Slide 12: A simple terrain site in southern France
Turbine | Turbine 3MW, operated by EDF EN Services Hub-height: 80 meters Rotor size: 90 meters
IEC mast | Top cup at hub-height 240 meters (2.7D) from turbine Free wind sector: [315°;335°] Temperature above 0.1°

## Slide 13: Wind speed : equivalent accuracy to 2-beam Wind Iris
Interpolated horizontal wind speed measurement is observed to achieve high accuracy and precision compared to IEC mast measurement.
No specific trend is observed (wind speed, wind direction,…)

## Slide 14: Uncertainty of power law interpolation
Hypothesis
Mast measurements were used to compute uncertainties due to power law profile assumptions
The uncertainty due to interpolation lies between 0.01 m/s and 0.08m/s at this site.
Reconstruction for a profile
Mean deviation between reconstructed profile
and measured wind speed

## Slide 15: Shear measurement compared with mast
Shear measured by 4-beam Wind Iris is observed to match mast measurements most of the time.
Some difference appears: due to difference in measurement heights?

## Slide 16: Turbulence intensity compared with mast
Turbulence intensity measured by 4-beam Wind Iris is observed to match the mast measurements most of the time.
Some difference appears: linked to atmosphere stability?

## Slide 17: Metrology of 4-beam Wind Iris in simple terrain
 | Accuracy | Precision
Horizontal Wind speed | 0.06 m/s | 0.18 m/s
Vertical wind shear | 0.014 | 0.092
Turbulence Intensity | -0.4 % (absolute) | 1.5 % (absolute)

## Slide 18: Validating 4-beam Wind Iris precision from power curve measurement
Power curve following the spirit of the IEC61400-12-1:2005 requirements (reduced wind sectors, temperature correction,…) is plotted with 4-beam Wind Iris measurements.
Mast based power curve is plotted over same time period so that devices are compared on same wind conditions.
Mean observed dispersion per bin is 81 kW for 4-beam Wind Iris and 111 kW for mast.
4-beam Wind Iris is observed to deliver a less scattered power curve than mast.
Notes: Mast
Lidar
IEC PC
Not reached
33 days
Ope. PC
58 days
8 days

## Slide 19: Power performance testing from 4-beam Wind Iris

## Slide 20: Power performance analysis
The 4-beam Wind Iris based power curve is in high agreement with the reference power curve -> no underperformance detected.
Nacelle anemometer wind measurement is biased (4% overestimation of the nacelle transfer function).
The nacelle anemometer based power curve is significantly under the reference power curve and erroneously shows underperformance.
Power curves are plotted following adapted requirements established within the EUDP project (Procedure for wind turbine power performance measurement with a two-beam nacelle lidar, Rozenn Wagner, Rebeca L Rivera, Ioannis Antoniou, Samuel Davoust, Troels F Pedersen, Mike Courtney, Ba-bak Diznabi , 2013)

## Slide 21: Shear effect on power performance
The higher the shear the more turbine performance deviates from reference (either underperforming or outperforming)

## Slide 22: Turbulence intensity effect on PP
The higher the TI the lower the performance of the turbine is.

## Slide 23: Conclusion

## Slide 24: Conclusions
Metrological performances of the new system are observed to be in continuity of the well established 2-beam Wind Iris.
4-beam Wind Iris provides both shear and turbulence intensity measurements with good accuracy and precision, although there are deviations which may be explained by the difference in measurement heights.
4-beam Wind Iris will allow for improvement in understanding of wind turbine performance in outer range conditions.